## Slide 1: Services RE3 & 4/1
Ian
5 Sept 2018

## Slide 2
Services to be installed in UXC during LS 2 , 2019 & 2020
HV cables and connectors, Umbilical cables USC to UXC. Coaxial cables from YE1PP to YE3 via mini-cable chains
LV cables & connectors on YE3
Fibre optics & Patch Panels from YE3 to YE1PP through to USC
Gas System piping, Impedances and Bulkheads on YE3
Cooling, flexibles on RE4 face of YE3, RE3/1 cooling comes with chambers

## Slide 3
HV System in UXC
Design change in the UXC area as the community has decided that the “3 pole “ is not a reliable solution.
This means that the connectors on the chambers and YE1 PP will be changed to what we have used in all tests since 2013. Namely the C.P.E. “Jupiter” coaxial connector.
HV order, There is a question on the insulator in the connectors. Is it Teflon and should we convert to PEEK. Technical documentation and EDH order ready No. 7367439
Resistances in boxes, simulating the chamber load, should be placed around the periphery to speed up commissioning.

## Slide 4
HV System in the USC
A new rack, S1H12, is installed for the RE3/1 and RE4/1 components.
The TDR version, an extension of the present system is planned.
However an alternative preferred version is planned for testing in 2019
Layout from Anton with calculation of components.
Which ever solution is finally taken does not influence the cabling from UXC.

## Slide 5
HV system routing from USC to UXC
Routing from the USC HV racks to the YE1 PPs
The Main Cable chains in X0 under CMS
Cable sections and lengths, including spares, to be evaluated with CMS Integ Office, S. Bally and Boki

## Slide 6
Modification of the YE1 PP to accept the alternative Jupiter connecter
Cable routing done by Integration office. Requires finalising
YE1 Patch Panel (PP)

## Slide 7
LV Cables UXC, Rack to chamber on YE3
LV tests done, different cable sections and stiffnesses.
Measured time response, is it what CAEN state, good enough ?
Do test with FEBs on chamber.
Supply of 3.5km of the cable (in P5) used for LB
Connectors at chamber end are more or less defined , “D-Sub’ with sense pins
2 cables per chamber but what of the connection in the CAEN module ? Choice of two.
Assembly can be done in Geneva EPI or Cavitech , Boudry, Neuchatel
As with the HV the load should be simulated by a resistance in place to
aid in rapid commissioning.
With the crates and modules in place the cabling is more efficient and quicker.
Coomunication cable fro USC to Uxc to be installed in main cable chains Jan 2019 ?

## Slide 8
Offers from CERN Logistics office
Power Sense
2 x 4mm2 plus 2 x 0.5mm2 Novacavi
5.0Euro/m 0.92Euro/m Draka but wrong colour ! GEM use black ?
With colour change require 12 weeks including August from Novacavi and 10-12 fro Draka inc Aug
Offer for combined Sense and Power never received. Renewed request on 4 Sept 2018, waiting.

## Slide 9
Christoph suggested a test transiting between 3A and 6.5A simulating the Firmware upload state.
Voltage drop at “Operational load” of 15W/FEB is <1.5V at the worst case of 30m. At 6.5A the worst case V drop is ~3V, that is on the limit of CAEN module.
Now that BT is finished the test should be done with the real FEBs. The real difference in current and transient should be studied.
Voltage drop on RPC LB cable |  |  |  |  |  |  |  | 
 |  |  |  | ian Crotty |  |  |  | 
 |  |  |  | 06-Aug-18 |  |  |  | 
 | Set point | Current | Power in Load | PS voltage | Voltage | Voltage Drop | V Drop/m | Worst case drop 30m
 |  |  |  |  | Load |  |  | 
 | [V] | [A] | [W] | [V] | [V] | [V] | [V] | [V]
Sense | 5.00 | 3.11 | 15.52 | 5.70 | 4.99 | -0.71 | -0.04797 | -1.44
Sense | 5.00 | 6.70 | 33.43 | 6.53 | 4.99 | -1.54 | -0.10405 | -3.12
No Sense | 5.00 | 5.33 | 21.21 | 5.21 | 3.98 | -1.23 | -0.08311 | -2.49
No Sense | 5.00 | 2.83 | 12.90 | 5.21 | 4.56 | -0.65 | -0.04392 | -1.32
Sense | 5.00 | 3.10 | 15.47 | 5.70 | 4.99 | -0.71 | -0.04797 | -1.44
Sense | 5.00 | 6.71 | 33.42 | 6.53 | 4.98 | -1.55 | -0.10473 | -3.14
TB cable 2 x 2.5mm2 plus sense 2 x |  |  |  |  |  |  |  | 
Cable length = |  | 14.80 | [m] |  |  |  |  | 
Max able length = |  | 30.00 | [m] |  |  |  |  | 
Done with Iuri and Irakli |  |  |  |  |  |  |  | 

## Slide 10
Response time of the CAEN module to the transition from low current (3A) to the high current (6-7A) condition
Sense Voltage at Load
Voltage at CAEN power terminals
Time of switching from 3 to 6.5A
Response of the PS to the sense input = 1ms
Thanks Iuri for your help in getting this view

## Slide 11
Fibre optics.
What technology can we install, frame contract with Blown fibres of “Trunk cables “ a la Polish trigger team.
Specification, see Behzad talks.
This week we will understand exactly what is the specification and within a few weeks the provider, order by end of October 2018.
Wolfram has stated yesterday that we can install what we want but it must be in the plan.
Fibres to installed in week 5 of 2019 !!
The order should leave CERN before the end of September.

## Slide 12: Schematic
1
2
3
4
5
6

## Slide 13: Specification of Every Cisco 10 gig SFP transceiver
Cisco SFP+ | Wavelength (nm) | Interface | Cable Type | Core Size (Microns) | Cable Distance
Cisco SFP-10G-T-S | / | RJ-45 | Cat6a/cat7 | / | 100m*
Cisco SFP-10G-SR-S | 850 | Dual LC/PC | MMF | 62.5(OM1) | 26m
Cisco SFP-10G-SR | 850 | Dual LC/PC |  | 62.5(OM1) | 33m
Cisco SFP-10G-SR-X | 850 | Dual LC/PC |  | 50(OM2) | 66m
 |  |  |  | 50(OM2) | 82m
 |  |  |  | 50(OM3) | 300m
 |  |  |  | 50(OM4) | 400m
Cisco SFP-10G-LRM | 1310 | Dual LC/PC | SMF | G.652(OS2) | 300m
 |  |  | MMF | 62.5(OM1) | 100m
 |  |  |  | 50(OM2) | 220m
 |  |  |  | 50(OM3) | 220m
Cisco SFP-10G-LR | 1310 | Dual LC/PC | SMF | G.652(OS2) | 10km
*30 meters via 10Gbps, 50 meters via 5Gbps and 2.5Gbps, 100 meters via 1000Mbps

## Slide 14: RE4(3)/1 chambers gas distribution schematic
Sagana brass union
reduction from 8mm to 6mm
Legris brass unions (6mm)
8 mm copper pipe
6 mm copper pipe
RE4(3)/3 chamber
RE4(3)/2 chamber
RE4(3)/1 chamber
Impedance box 6 channel
Gas Rack
Bottom layer
Top layer
Supply
Return
Parallel flow

## Slide 15: From pre-distribution to distribution UXC racks
We need:
  Supply: 2 inox pipes of Ø 16 (90m for RE+ and 77m for RE- ~170m).
  Return: 2 inox pipes of Ø 22 (90m for RE+ and 77m for RE- ~170m).
  4 Flexible pipes for the chains (20 – 30 m of 2 x Ø 16 and 2 x Ø 22).
From USC gas room to UXC gas racks: new pipes will be installed (2 pipes per disk (negative/positive) for supply/return lines).
To be installed Jan 2019.

## Slide 16
YE3 gas rack to be modified to make the 12 channels independent from the present channels
Impedance boxes positioned at three points around the periphery for both RE4/1 and RE3/1
Impedance box for parallel operation, 2 channel shown

## Slide 17
Gas
Piping (8x6mm) from YE3 rack to bulkheads on the peripheral “green structure’ area.
Require spread sheet of take off points on the top of the rack to the bulkheads, and so chambers.
Require detailed tech drawings of this piping.
Integrate the 6 impedance boxes around the periphery, naming tags (metal), fabricated and tested then installed and commissioned.
Lay piping (6x4mm ?) ,with labels , from the periphery to the chamber position and place short cut to commission system.
Calculate impedances based on ‘higher” flow rates, as shown in GIF tests, and so define the final piping to the chambers from the periphery

## Slide 18
Cooling
RE3/1
Verify the assembly of the flexibles with simplified safer version, namely direct from RE3/2 PP to RE3/1 PP with three elements only, as opposed to 3 times as many unions with a “L” Cu pipe producing a lever arm on the chamber PP.
RE4/1
Modify the present Mini-manifolds to take additional “take offs” direct to the chambers with flexibles all the way to the RE4/1 PP.

## Slide 19
Mini manifold . To be modified to accept the necessary “T”s for the RE4/1 cooling.
The same flow regulators will be kept with individual cut off valves to each chamber, if costs permit.

## Slide 20
This list will be ordered, simplified , from the USA or from XX in P5 in inch size.

## Slide 21: Man power

## Slide 22
RE3/1 & RE4/1 Service Installation during LS2

## Slide 23
RE4 interventions during LS2

## Slide 24: Contacts and action
HV. Outstanding points with CPE, assembly QC methods and timeline
LV
  Do test with New FEB while loading firmware and max cable length ~30m
  S. Bally for exact cable lengths
  Contact Houari for cable length in P5 store.
Fibres, check where we are with spec and supplier in next two weeks as Bezhad and Zhenan are here.
  Check procurement details with Austin.
Gas see with PK and Anton/Alexander/ZEC
See with cooling suppliers, French and XX at P5 for the inch size on “T” mini manifolds and the metric on the chamber.

## Slide 25: Additional Material

## Slide 26
https://www.epi.ge.ch/prestations-entreprises-particuliers/ateliers/microtechnique/

## Slide 27: RE4(3)/1 chambers gas distribution schematic
Sagana brass union
reduction from 8mm to 6mm
Legris brass unions (6mm)
8 mm copper pipe
6 mm copper pipe
RE4(3)/3 chamber
RE4(3)/2 chamber
RE4(3)/1 chamber
Capillary
Gas Rack
Bottom layer
Top layer
Supply
Return
Parallel flow

## Slide 28: RPC Power System Review
Anton Dimitrov, Ian Crotty
RPC Upgrade Mini Workshop, 05.Sep.2018, CERN

## Slide 29
The HV power system will be based on CAEN EASY system
(practically, extension of the existing HV system):
SY4527 – 1 (existing)
A3485S - 2 modules (existing)
4 EASY3000 crates
A3512N – 12 HV Boards
4 HV multiconductor cables per endcap. Each with 20 HV channels
All must go via the YE1’s main cable chains (MCC)
Long term maintenance will be secured by CERN-CAEN contract.
To be integrated in the RPC DCS.
A3512N
A3485S
SY4527
EASY3000
HV Power System Upgrade for RE3/1 & RE4/1 (I)
USC
UXC
TDR Baseline: Variant I
1

## Slide 30
A3485S
A3485S
A3485S
A3485S
SY4527
SY4527
S1H07
HEX
HEX
ENDCAP RACKS S1H11 – S1H08
Crate4
RE-3/1 RE-4/1
Crate4
RE+3/1 RE+4/1
HV Distribution Box
HV Distribution Box
HV Distribution Box
HV Distribution Box
HV Distribution Box
HV Distribution Box
HEX
HEX
S1H12
HV Distribution Box
HV Distribution Box
48V DC cable connected after Crate3 in rack S1H09
48V DC cable connected after Crate3 in rack S1H11
1 new rack (already defined S1H12, to be equipped)
2 EASY crates
8 umbilical cables all arriving to one rack
6 A3512N HV boards in each Crate (12 boards in total)
8 HV Distributor Boxes in total (all in one rack)
2

## Slide 31
The HV power system will be based on CAEN NON-EASY system:
SY4527 – 2 (new)
New HV boards – prototype ready in 2019
4 HV multiconductor cables per endcap. Each with 20 HV channels
All must go via the YE1’s main cable chains (MCC)
Long term maintenance will be secured by CERN-CAEN contract.
To be integrated in the RPC DCS.
A3512N
A3485S
SY4527
EASY3000
HV Power System Upgrade for RE3/1 & RE4/1 (II)
USC
UXC
Preferred Variant: Variant II
3

## Slide 32
A3485S
A3485S
A3485S
A3485S
SY4527
SY4527
S1H07
HEX
HEX
ENDCAP RACKS S1H11 – S1H08
HV Distribution Box
HEX
HEX
S1H12
- No EASY system  no EASY crates
- 2 SY4527 mainframes more
- 8 umbilical cables all arriving to one rack
- New type of HV boards!
- 8 HV Distributor Boxes in total (all in one rack), but bigger in size since all inputs and outputs should be at the back side of the rack
- No rack cross reference in DSS Action Matrix
SY4527
SY4527
HV Distribution Box
HV Distribution Box
HV Distribution Box
HV Distribution Box
HV Distribution Box
HV Distribution Box
HV Distribution Box
4

## Slide 33
24 A3016 LV boards in total to power the FE of all RE3/1 and RE4/1 chambers.
9U in each YE3 X4Far tower rack for
  1 CAEN EASY3000S crate
  1 CAEN A3486S MAO- AC/DC 48V PS
15U in each X2Near YE3 tower rack for
  2 CAEN EASY3000S crates
  1 CAEN A3486S MAO- AC/DC 48V PS
SY4527 with A1676A
CAEN EASY3000S
with CAEN A3016 in X2Near && X4Far racks
A3486S
48V Service power
50p Communication bus
USC
UXC
YE1 LV PP
RPC LV Power System for RE3/1 & RE4/1
1
going directly to YE3
Relay
YE+3: X2J52, X3J51, X3A51, X4A51
YE-3: X2V52, X3V51,X3S51, X4S51
Existing RE3&RE4 detector

## Slide 34
mini Cable Chains for OF
LV Easy crates
Near Side
Far Side
mini Cable Chains for OF
LV Easy crates
X2
X4
RPC LV Services
2
X2Near – X4Far diagonal

## Slide 35
Air Deflector
HEX
X2J52
RE3 LBB
Cable Distribution
RE A3486S AC/DC-0 CH0 CH1
EMU Maraton
EMU Maraton
Turbine
Power Distribution
X3J51
X4A51
YE+3
Near Side
Far Side
RE4 LBB
HEX
HEX
RE3/1 FE EASY3000S Crate0
RE4/1 FE EASY3000S Crate1
HEX
Turbine
Power Distribution
EMU
EMU VME
Air Deflector
HEX
RE3 LBB
Cable Distribution
RE3-4 FEB EASY3000S Crate2
RE3-4 LBB EASY3000S Crate3
RE A3486S AC/DC-1 CH0 CH1
RE4 LBB
Power Dist.
Power Dist.
Power Dist.
HEX
Turbine
Power Distribution
RE3 LBB
RE4 LBB
EMU Maraton
HEX
EMU Maraton
HEX
Air Deflector
Alignment DAQ
RE3-4/1 FE EASY3000S Crate2
RE A3486S AC/DC-1 CH0 CH1
Power Dist.
HEX
Turbine
Power Distribution
HEX
Air Deflector
RE4 LBB
RE4 LBB
Cable Distribution
RE3-4 LBB EASY3000S Crate1
RE3-4 FEB EASY3000S Crate0
RE A3486S AC/DC-0 CH0 CH1
X3A51
EMU
EMU VME

## Slide 36
DCS daisy chain
YE+3
Near Side
Far Side
X2J52
X3J51
X3A51
X4A51
YE3Near and YE3far equipment (including existing RE3&RE4 detector on X3) on separate branch controllers, therefore 4 new A1676 branches will be needed in total.