RPC Endcap		Station 1

- HV Services

| RPC | Nos Chamber Station | Channel /chamber | Conductor | cables | Cables /20 | Section /20 (cm2) | Section /Station (cm2) | Remote Distrib. Board /60 | Cables Leaving Yoke to SY1527 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| RE1/1 | 36 | 3 | 9 | 1 | 2 | 3.54 | 63.7 |  |  |
| RE1/2 | 36 | 3 | 9 | 1 | 2 | 3.54 | 63.7 |  |  |
| RE1/3 | 36 | 3 | 9 | 1 | 2 | 3.54 | 63.7 |  |  |
| Total | 108 | 9 | 27 | 3 | 6 | 10.62 |  | 3 |  |
| Total/ ## Station/ Endcap |  |  |  | 108 | 108 |  | 191.16 | 18 | Section 25.92=cm2 |

# HV Cable has an estimated diameter of  13.3mm, instead of the 16.2mm

specified for the Barrel ( that is 9 conductors of 3mm).

8 Channels to each RDB (Remote Distribution Board). Dimensions 266mm x 80mm x 380mm

Note there is 1 spare RDB output/20 ie 3 spare gas gap supplies.

500V supply cable from SY1527 to RDB is 3x1mm2 (diam. 12mm)

eg; SCEM 04.08.61.525.2 as a first estimate.

Cable section is 2 in cm2.                                                                                                    IMC  2000.06