## Slide 1: L/HVPS CAEN Contracts
fv/19May2016

## Slide 2: CAEN P&M Contracts
B1226 P&M
2009
2010
2011
2012
2019
B1281 M
B1291 P
B1226 P
B1337 P
2013
Stopped by CAEN
Addendum list
Re-launched 2012
LHC equipment list
As B1281
LHC equipment list
Revisit situation once experiments have clearer view
of LS2 needs
2 price lists
LHC & Addendum list
P=Procurement
M=Maintenance
End 31.12.2019
End 31.12.2019
Continue post LS1
B1387 P
Addendum list
2014
Spent 5.1M till Dec 2015
Spent 1.8M till Dec 2015
2015
B1451 P
Addendum list
2016

## Slide 3: CAEN contracts: B-1173 (FC/4832-obsolete) + B-1226 (IT-3330)+ B-1281 (FC/5277) + B-1291 (obsolete)+ B-1337 (obsolete) + B1387 (obsolete) (in CHF)
Chart: 
| TOTAL | B1173 | B1226 | B1281 | B1291 | B1337 | B1387
2001
2002 | 749815 | 749815
2003 | 916796 | 916796
2004 | 970986 | 970986
2005 | 1019606 | 1002405 | 17201
2006 | 2027460 | 1865322 | 162138
2007 | 3264690 | 1365779 | 1898911
2008 | 1638963 | 463803 | 1175160
2009 | 239058 | 56812 | 4158 | 178088
2010 | 861326 | 27944 | 361698 | 471684
2011 | 830909 | 25973 | 283782 | 521154
Column2

## Slide 4
Chart: Annual failure rate
| 2008 | 2009 | 2010 | 2011 | 2012 | 2013 | 2014
Wiener | 0.036573 | 0.027622 | 0.011692 | 0.012785 | 0.015755 | 0.019447 | 0.023512
Caen | 0.019083 | 0.06687 | 0.045863 | 0.033515 | 0.029506 | 0.021022 | 0.032853
Annual fault reports statistics
CAEN B1281 Maintenance Contract
On-Site maintenance:
All LHC experiments
75’000 CHF per year (54’000 CHF supported by CMS)
Plus software maintenance fee

## Slide 5
2014 repair and preventive maintenance statistics
Chart: Wiener
| 41640 | 41671 | 41699 | 41730 | 41760 | 41791 | 41821 | 41852 | 41883 | 41913 | 41944 | 41974
Repair requests | 26 | 13 | 1 | 10 | 11 | 3 | 9 | 12 | 20 | 18 | 10 | 10
Preventive maintenance | 58 | 13 | 10 | 12 | 8 | 24 | 4 | 2 | 0 | 0 | 13 | 0
Chart: Caen
| 41640 | 41671 | 41699 | 41730 | 41760 | 41791 | 41821 | 41852 | 41883 | 41913 | 41944 | 41974
Repair requests | 8 | 20 | 12 | 10 | 10 | 12 | 17 | 17 | 14 | 37 | 42 | 13
Preventive maintenance | 5
CAEN B1281 Maintenance Contract

## Slide 6
Caen - Repair turnaround time (at contractor’s premises)
Between ‘repair authorized’ and ‘returned from repair’ actions
Caen 2014
Average turnaround time: 29.6 days
(In 2013: 51.8 days)
On-Site maintenance:
All LHC experiments
75’000 CHF per year (54’000 CHF supported by CMS)
Plus software maintenance fee

## Slide 7
CAEN contracts up-to-now
Running smoothly
Appreciated by users
Stable failure rate since 2010 (~3%, no sign of wearout)
Acceptable turn-around time (avg 30 days)
Stock of critical components available
~Monthly visits and meetings
Costly
Probably no change till 2019
Chart: Annual failure rate
| 2008 | 2009 | 2010 | 2011 | 2012 | 2013 | 2014
Wiener | 0.036573 | 0.027622 | 0.011692 | 0.012785 | 0.015755 | 0.019447 | 0.023512
Caen | 0.019083 | 0.06687 | 0.045863 | 0.033515 | 0.029506 | 0.021022 | 0.032853
Annual fault reports statistics

## Slide 8
CAEN contracts after 2019
Extension to LS3 likely
  2019-2023/24?
  Good service to us
  Good money to CAEN
  No sign of equipment wearout (yet)
  On-site service is expensive (but convenient)
  Subject to re-evaluation of spares situation
  Subject to re-negociation of prices
  Subject to agreement with other LHC experiments
Beyond LS3
  Existing equipment will wear-out
    Not well matched to long-term contract
  CAEN has a tradition of repairing aged equipment if they can
    Best effort
  New equipment subject to new Tender
    Maintenance contract linked to new equipment